Changes in original posejdon design  (https://github.com/pachterlab/poseidon/tree/release)

BASE

- All socket head screws were swapped for M3x16 (The holes in which the screws are mounted were deepend accordingly)

- M6 threaded rod changed to TR8x2 screw (8mm pitch)

- The hole below the syringe (Where the lead screw is supposed to go) was rounded and the diameter was changed from 3.5mm to 6.5mm

- The syringe holder module was changed entirely (both the knob and the holder itself) and the elastic foam was added in order to ensure stability of the syringe.

- Knob for M5 comes from Thingiverse project (Compact M5 Nut Knob

by B_EMME3D https://www.thingiverse.com/thing:774472 )

CARRIAGE

- The crosswise hole in the middle of the carriage as well as on the bottom of it were sealed shut.

- The hole allowing the lead screw to pass through the carriage was enlarged and deepend to fit the TR8x2 nut. 4 additional holes were made for M3x10 screws to hold the nut in place

- The cavities made for the linear bearings were shortened from 20mm to 19.1mm (0.05mm free space on both sides of the bearings)

- The back of the carriage was drawn-out to fit the T8x2 nut

- Pump cover slide element was removed from the project

SOFTWARE

- steps2mm and mm2steps functions were changed to match the screw pitch of the T8x2 lead screw. One revolution of the motor shaft equals 8mm displacement of the carriage.

- The area of the syringe was changed accordingly to the used model (It is important to establish the area of the model you will be using in your project beforehand).

- The title of the GUI was changed to "Poseidon Pumps Controller - Biomarkers Analysis Lab AGH 2024"

Every part in the project was printed out with PETG 3 Perimeters and 25% infill.

More information regarding the specific changes from the original design can be found either in the BOM or the 3D models themselves.